GREIT:  Consensus EIT algorithm for lung images
Validation and parameter selection for GREIT
Andy Adler, John Arnold, Richard Bayford, Andrea Borsic, Brian Brown, Paul Dixon, Theo J.C. Faes, Inéz Frerichs, Hervé Gagnon, Yvo Gärber, Bartlomiej Grychtol, Günter Hahn, William R B Lionheart, Anjum Malik, Janet Stocks, Andrew Tizzard, Norbert Weiler, Gerhard Wolf
GREIT: Graz Consensus Reconstruction Algorithm for EIT

---
What problem are we trying to solve?
What problem are we trying to solve?
EIT useful to image regional lung air flow. But, we can't compare regions unless
  - Consistent response:
    - Air volume in two
      places must show
      the same image
  - Accurate positions
  - No strange shapes
R1
R2
R3
R4
Our main goal is consistency

---
Requirements
Amplitude
{} prop sum_k [hat{bold x}]_k
Set ¼-Max
Amplitude ( )
Reconstructed
Image ( )
{hat{bold x}}_q
hat{bold x}
Position Error
r_t - r_q
rt
rq
A_q
Resolution
sqrt{A_q / A_0}
Simulation
Target
Shape
Deformation
A_{outside} / A_q
Ringing
Inverted image

---
Consensus Ranking
Amplitude
#1. Requirement: uniform
#2. Requirement: uniform and small
#4. Requirement: uniform
#6. Requirement: small
#5. Requirement: small
#3. Requirement: small
{} prop sum_k [hat{bold x}]_k
Position Error
r_t - r_q
Resolution
sqrt{A_q / A_0}
Shape
Deformation
A_{outside} / A_q
Ringing

---
Formulation: Training data
GR
Simulation
Target
Desired image
- based on
 requirements
Reconstructed
image
- to be optimized
Amplitude
{bold x}_t^{(k)} {hat{bold x}}^{(k)} {tilde{bold x}}^{(k)}

---
Type of Signal	Training Inputs (measurements)	Desired Output (reconstructed images)
Conductivity targets
Noise / - electronic noise / - electrode movement
Circular
PSF
{bold y}_t^{(k)}
{tilde {bold x}_t}^{(k)}
0 outside
PSF
k+1
k
k–1
{tilde {bold x}_n}^{(k)}=0
{bold y}_n^{(k)}
Desired image for noise
input is zero
Minimize:
GREIT Reconstruction Matrix
[Speaker Notes: \begin{align*} \epsilon^2 = \bigg\| &[\cdots\tilde{\bf x}_t^{(k)} \cdots | \cdots {\bf 0} \cdots] - \\ {\bf R}&[\cdots\tilde{\bf y}_t^{(k)} \cdots | \cdots \tilde{\bf y}_n^{(k)} \cdots] \bigg\|^2_{\bf W} \end{align*}]

---
Example Images
Backproj	NOSER	GREIT
 i

---
Performance Measures
Amplitude Response (AR) Position Error (PE)
GREIT NOSER Backproj

---
GREIT Variants
Body shape
Lung
conductivity
Fat thickness
Image from 6-7th intercostal
space: male visual human
How important?
How many models?

---
Example: reference conductivity
Backproj
GREIT
Lung σ
 % of other
   tissue
 100%
    30%
3D model
2D model
Normalized
Difference
Lung injured piglet at 10cmH2O. Data from Frerichs et al (2003)

---
“Roadmap”
“Roadmap”
Step 1: Agree on “ingredients” and “requirements”
Step 2: Develop an algorithm framework
  - Physiol. Meas. 30:S35-S55, 2009
Step 3: Distribute tools
  - Publish software (EIDORS 3.4)
  - Analysis tools (including online)
Step 4: Validate
  - Experience using GREIT
  - Systematic evaluation with “real” data
June 09
Oct. 09

---
Calibrated animal tests
Calibrated animal tests
Adler et al (1996)
Frerichs et al (2003)
Hahn et al (2006)
To Do: Design tests
based on these data

---
Clinical case
Clinical data
- Patient data male 59 yrs 188 cm 120 kg
- Current diagnosis Sepsis with acute lung injury Acute renal failure (continuous dialysis) Atelectasis left lower lung lobe
- Medical history Implantation of cardiac pacemaker Arterial hypertension
- EIT measurements performed in the ICU
- Mode Continuous positive airway pressure ventilation with assisted spontaneous breathing (CPAP/ASB)
- FIO2 0.5 PEEP 9 cmH2O Frequency 25 breaths/min Minute ventilation 15.1 l/min
- During the EIT measurement of 180 s duration approx. after 60 s PEEP was reduced from 9 to 5 cmH2O and after 120 s increased to 13 cmH2O.
Ppeak 20 cmH2O Pmean 13 cmH2O at PEEP 9 cmH2O SO2 97 %
Ppeak 16 cmH2O Pmean 9 cmH2O at PEEP 5 cmH2O SO2 92 %
Ppeak 24 cmH2O Pmean 16 cmH2O at PEEP13 cmH2O SO2 97 %
To Do: Design tests
based on these and other clinical data